## Sheet1
Sample ID | Sample Rack ID | Sample Rack Position | APU | APU Location | Test Type | Specimen Type | Operator ID at time of Order | Operator ID at time of Release | Assay Name | Assay Number | Assay Version | Final Result Value | Secondary Result | IC Response Value | Cellular Control Response | Max Ratio | Interpretation | Result Flags | Result Code | Dilution Name | Instrument Response Value | Date of Completion | Result Status | Transmission Status | Transmission Date / Time | Retrieve Date / Time | Archive Status | Archive Date/Time | System Serial Number | System Software Version | Message Code | Amp Kit Lot Number | Amp Kit Unit Number | Amp Kit Lot Expiration Date | Calibrator Lot Number | Calibration Date Time | Calibrator Lot Expiration Date | Sample Onboard Time | Initiation Date Time | Sample Prep Lot Number | Sample Prep Lot Expiration Date | Sample Prep Unit Number | Lysis Lot Number | Lysis Lot Expiration Date | Ethanol Lot Number | Ethanol Expiration Date | Vapor Barrier Lot Number | Vapor Barrier Lot Expiration Date | Diluent Lot Number | Diluent Lot Expiration Date | Auxiliary Sample Prep Lot Number | Auxiliary Sample Prep Unit Number(s) | Auxiliary Sample Prep Lot Expiration date | IRU Lot Number | Tip 1000 Lot Number | Tip 50 Lot Number | Comment | Control Master Set Lot | Control Master Set Lot Exp Date | Neg1 Control Lot Number | Neg1 Control Lot Expiration Date | Neg1 Sample ID | Neg1 Rack ID | Neg1 Position | Neg1 Result Value with Units | Neg1 Result Flags | Neg1 Message Code | Neg1 Date and Time of Completion | Pos1 Control Lot Number | Pos1 Control Lot Expiration Date | Pos1 Sample ID | Pos1 Rack ID | Pos1 Position | Pos1 Result Value with Units | Pos1 Result Flags | Pos1 Message Code | Pos1 Date and Time of Completion | Pos2 Control Lot Number | Pos2 Control Lot Expiration Date | Pos2 Sample ID | Pos2 Rack ID | Pos2 Position | Pos2 Result Value with Units | Pos2 Result Flags | Pos2 Message Code | Pos2 Date and Time of Completion | Assay / Target1 Name | Assay / Target2 Name | Assay / Target3 Name | Assay / Target4 Name | Assay / Target5 Name | Assay / Target6 Name | Target 1 Target Response | Target 2 Target Response | Target 3 Target Response | Target 4 Target Response | Target 5 Target Response | Target 6 Target Response | Target 1 Max Ratio | Target 2 Max Ratio | Target 3 Max Ratio | Target 4 Max Ratio | Target 5 Max Ratio | Target 6 Max Ratio | Target 1 Reported Result | Target 2 Reported Result | Target 3 Reported Result | Target 4 Reported Result | Target 5 Reported Result | Target 6 Reported Result | Target 1 Interpretation | Target 2 Interpretation | Target 3 Interpretation | Target 4 Interpretation | Target 5 Interpretation | Target 6 Interpretation
2506640801 | A145644 | 4 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.45 |  | 0.242 | Target Not Detected |  |  | UNDILUTED | 29.85 | 07.01.2025 11:17 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:23 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640802 | A145644 | 5 | 1 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 09:23 PM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 5011 - Pipettor sample aspiration error type: (Clot Detected), pipettor: (2) | 399437 | 01317,01545 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:04 PM | 07.01.2025 09:23 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 | [36 empty] | HIV-DBS
2506640803 | A145644 | 6 | 1 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.99 |  | 0.259 | Target Not Detected |  |  | UNDILUTED | 29.05 | 07.01.2025 11:17 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:23 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640804 | A145644 | 7 | 1 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.85 |  | 0.232 | Target Not Detected |  |  | UNDILUTED | 30.23 | 07.01.2025 11:18 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:23 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640805 | A145644 | 8 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.94 |  | 0.339 | Target Not Detected |  |  | UNDILUTED | 29.23 | 07.01.2025 11:18 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:23 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640806 | A145644 | 9 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.65 |  | 0.326 | Target Not Detected |  |  | UNDILUTED | 29.30 | 07.01.2025 11:18 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:23 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640807 | A145644 | 10 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.65 |  | 0.294 | Target Not Detected |  |  | UNDILUTED | 28.00 | 07.01.2025 11:18 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:23 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640808 | A145644 | 11 | 1 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.20 |  | 0.008 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:18 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:23 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640809 | A145644 | 12 | 2 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.55 |  | 0.289 | Target Not Detected |  |  | UNDILUTED | 29.62 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640810 | A096567 | 1 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.29 |  | 0.006 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:51 AM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 01532,01526 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:15 PM | 07.01.2025 10:58 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506640811 | A096567 | 2 | 2 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.62 |  | 0.274 | Target Not Detected |  |  | UNDILUTED | 27.49 | 07.02.2025 12:51 AM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 01532,01526 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:15 PM | 07.01.2025 10:58 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640812 | A096567 | 3 | 2 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.65 |  | 0.201 | Target Not Detected |  |  | UNDILUTED | 28.37 | 07.02.2025 12:51 AM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 01532,01526 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:15 PM | 07.01.2025 10:58 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640813 | A096567 | 4 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.22 |  | 0.008 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:51 AM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 01532,01526 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:15 PM | 07.01.2025 10:58 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640814 | A096567 | 5 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.01 |  | 0.007 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:51 AM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 01532,01526 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:15 PM | 07.01.2025 10:58 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506640815 | A096567 | 6 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.44 |  | 0.252 |  |  |  | UNDILUTED | 29.47 | 07.02.2025 12:51 AM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 01532,01526 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:15 PM | 07.01.2025 10:58 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506640816 | A096567 | 7 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.32 |  | 0.007 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:51 AM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 01532,01526 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:15 PM | 07.01.2025 10:58 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506640817 | A096569 | 12 | 4 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.96 |  | 0.288 |  |  |  | UNDILUTED | 28.16 | 07.02.2025 12:04 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506640818 | A075770 | 1 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 4,153 Copies/mL |  | 15.82 |  | 0.394 |  |  |  | UNDILUTED | 22.97 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640819 | A075770 | 2 | 2 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.11 |  | 0.262 | Target Not Detected |  |  | UNDILUTED | 30.07 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640820 | A075770 | 3 | 2 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 14.95 |  | 0.216 | Target Not Detected |  |  | UNDILUTED | 28.81 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640901 | A075770 | 4 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.82 |  | 0.009 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640902 | A075770 | 5 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.58 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640903 | A075770 | 6 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.58 |  | 0.289 | Target Not Detected |  |  | UNDILUTED | 29.80 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640904 | A075770 | 7 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.64 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640905 | A075770 | 8 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.84 |  | 0.358 | Target Not Detected |  |  | UNDILUTED | 27.69 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640906 | A075770 | 9 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.66 |  | 0.232 | Target Not Detected |  |  | UNDILUTED | 29.52 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640907 | A075770 | 10 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.61 |  | 0.244 | Target Not Detected |  |  | UNDILUTED | 28.96 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640908 | A075770 | 11 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.08 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:33 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:39 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640909 | A075770 | 12 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.91 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:49 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:55 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640910 | A145644 | 1 | 1 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.68 |  | 0.260 | Target Not Detected |  |  | UNDILUTED | 27.72 | 07.01.2025 11:17 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:23 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640911 | A145644 | 2 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.13 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:17 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:23 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640912 | A145644 | 3 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.21 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:17 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:04 PM | 07.01.2025 09:23 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640913 | A145645 | 8 | 4 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.89 |  | 0.398 | Target Not Detected |  |  | UNDILUTED | 27.52 | 07.02.2025 12:05 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640914 | A145645 | 9 | 4 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.80 |  | 0.327 |  |  |  | UNDILUTED | 27.71 | 07.02.2025 12:05 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506640915 | A145645 | 10 | 4 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.80 |  | 0.274 | Target Not Detected |  |  | UNDILUTED | 28.78 | 07.02.2025 12:05 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640916 | A145645 | 11 | 4 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.98 |  | 0.274 | Target Not Detected |  |  | UNDILUTED | 28.46 | 07.02.2025 12:05 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640917 | A145645 | 12 | 1 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.33 |  | 0.259 | Target Not Detected |  |  | UNDILUTED | 29.28 | 07.02.2025 12:20 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640918 | A096569 | 1 | 3 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.34 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:49 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 09:56 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506640919 | A096569 | 2 | 3 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.74 |  | 0.002 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 11:49 PM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 09:56 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506640920 | A096569 | 3 | 3 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.71 |  | 0.227 | Target Not Detected |  |  | UNDILUTED | 29.88 | 07.01.2025 11:49 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 09:56 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641001 | A096569 | 4 | 3 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.92 |  | 0.007 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:49 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 09:56 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641002 | A096569 | 5 | 3 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.04 |  | 0.278 | Target Not Detected |  |  | UNDILUTED | 29.39 | 07.01.2025 11:49 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 09:56 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641003 | A096569 | 6 | 3 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.72 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:49 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 09:56 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641004 | A096569 | 7 | 3 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.54 |  | 0.098 | Target Not Detected |  |  | UNDILUTED | 32.88 | 07.01.2025 11:49 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 09:56 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641005 | A096569 | 8 | 3 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.19 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:49 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 09:56 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641006 | A096569 | 9 | 3 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.86 |  | 0.236 | Target Not Detected |  |  | UNDILUTED | 27.55 | 07.01.2025 11:49 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 09:56 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641007 | A096569 | 10 | 3 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.98 |  | 0.282 | Target Not Detected |  |  | UNDILUTED | 29.59 | 07.01.2025 11:49 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 09:56 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641008 | A096569 | 11 | 3 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.13 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:49 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 09:56 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641009 | A076010 | 4 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.56 |  | 0.233 |  |  |  | UNDILUTED | 29.73 | 07.02.2025 12:21 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641010 | A076010 | 5 | 1 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.29 |  | 0.010 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:21 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641011 | A076010 | 6 | 1 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.01 |  | 0.277 | Target Not Detected |  |  | UNDILUTED | 29.94 | 07.02.2025 12:21 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641012 | A076010 | 7 | 1 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.71 |  | 0.229 | Target Not Detected |  |  | UNDILUTED | 28.72 | 07.02.2025 12:21 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641013 | A076010 | 8 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.43 |  | 0.344 | Target Not Detected |  |  | UNDILUTED | 28.15 | 07.02.2025 12:21 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641014 | A076010 | 9 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.09 |  | 0.007 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:21 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641015 | A076010 | 10 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.32 |  | 0.001 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:21 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641016 | A076010 | 11 | 1 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.10 |  | 0.269 | Target Not Detected |  |  | UNDILUTED | 28.09 | 07.02.2025 12:21 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641017 | A076010 | 12 | 2 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.10 |  | 0.296 | Target Not Detected |  |  | UNDILUTED | 27.19 | 07.02.2025 12:36 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641018 | A145645 | 1 | 4 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.20 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:05 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641019 | A145645 | 2 | 4 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.71 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:05 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641020 | A145645 | 3 | 4 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.49 |  | 0.308 | Target Not Detected |  |  | UNDILUTED | 27.75 | 07.02.2025 12:05 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641101 | A145645 | 4 | 4 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.85 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:05 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641102 | A145645 | 5 | 4 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.57 |  | 0.359 | Target Not Detected |  |  | UNDILUTED | 26.67 | 07.02.2025 12:05 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641103 | A145645 | 6 | 4 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.23 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:05 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641104 | A145645 | 7 | 4 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.61 |  | 0.249 | Target Not Detected |  |  | UNDILUTED | 30.01 | 07.02.2025 12:05 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:05 PM | 07.01.2025 10:11 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641105 | A076045 | 12 | 4 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.65 |  | 0.297 |  |  |  | UNDILUTED | 27.89 | 07.02.2025 01:08 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 11:15 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641106 | A096575 | 1 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.67 |  | 0.001 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:36 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641107 | A096575 | 2 | 2 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 2,141 Copies/mL |  | 15.63 |  | 0.323 |  |  |  | UNDILUTED | 23.64 | 07.02.2025 12:36 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641108 | A096575 | 3 | 2 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.42 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:36 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641109 | A096575 | 4 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.39 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:37 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641110 | A096575 | 5 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.40 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:37 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641111 | A096575 | 6 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.39 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:37 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641112 | A096575 | 7 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.87 |  | 0.211 | Target Not Detected |  |  | UNDILUTED | 29.42 | 07.02.2025 12:37 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641113 | A096575 | 8 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.80 |  | 0.005 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:37 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641114 | A096575 | 9 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 18.03 |  | 0.007 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:37 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641115 | A096575 | 10 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.87 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:37 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641116 | A096575 | 11 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 |  |  | -1.00 |  | 2.466 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:37 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 1998 - Fluorescence signal evaluation did not meet expected criteria. Excessive noise was detected in the reference signal. | 399437 | 01318,01520 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:06 PM | 07.01.2025 10:43 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641117 | A096575 | 12 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.54 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:52 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641118 | A076010 | 1 | 1 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.16 |  | 0.008 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:20 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641119 | A076010 | 2 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.91 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:21 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641120 | A076010 | 3 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.68 |  | 0.001 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:21 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:06 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641201 | A075768 | 8 | 4 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 18.56 |  | 0.004 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:09 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 11:15 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641202 | A075768 | 9 | 4 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.91 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:09 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 11:15 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641203 | A075768 | 10 | 4 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.00 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:09 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 11:15 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641204 | A075768 | 11 | 4 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.10 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:09 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 11:15 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641205 | A075768 | 12 | 1 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.32 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:24 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641206 | A076045 | 1 | 3 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 18.17 |  | 0.281 |  |  |  | UNDILUTED | 28.62 | 07.02.2025 12:52 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641207 | A076045 | 2 | 3 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.63 |  | 0.251 |  |  |  | UNDILUTED | 30.33 | 07.02.2025 12:52 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641208 | A076045 | 3 | 3 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.26 |  | 0.311 | Target Not Detected |  |  | UNDILUTED | 27.75 | 07.02.2025 12:52 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01318,01520 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641209 | A076045 | 4 | 3 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.75 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:52 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641210 | A076045 | 5 | 3 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.91 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:52 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641211 | A076045 | 6 | 3 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.11 |  | 0.329 | Target Not Detected |  |  | UNDILUTED | 27.76 | 07.02.2025 12:53 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641212 | A076045 | 7 | 3 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.23 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:53 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641213 | A076045 | 8 | 3 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.75 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:53 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641214 | A076045 | 9 | 3 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.25 |  | 0.269 | Target Not Detected |  |  | UNDILUTED | 27.48 | 07.02.2025 12:53 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641215 | A076045 | 10 | 3 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 18.15 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:53 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641216 | A076045 | 11 | 3 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 18.64 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:53 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 10:59 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641217 | A075768 | 4 | 4 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.03 |  | 0.005 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:08 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 11:15 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641218 | A075768 | 5 | 4 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 18.19 |  | 0.004 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:08 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 11:15 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506641219 | A075768 | 6 | 4 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.38 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:08 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 11:15 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641220 | A075768 | 7 | 4 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.84 |  | 0.266 | Target Not Detected |  |  | UNDILUTED | 27.64 | 07.02.2025 01:09 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:07 PM | 07.01.2025 11:15 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641301 | A096998 | 8 | 4 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.57 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:17 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641302 | A096998 | 9 | 4 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.99 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:17 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641303 | A096998 | 10 | 1 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.75 |  | 0.256 |  |  |  | UNDILUTED | 28.54 | 07.02.2025 02:26 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:31 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506641304 | A096998 | 11 | 1 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.26 |  | 0.246 |  |  |  | UNDILUTED | 29.16 | 07.02.2025 02:26 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:31 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506641305 | A096998 | 12 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.86 |  | 0.290 | Target Not Detected |  |  | UNDILUTED | 26.78 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:31 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641306 | A096996 | 1 | 3 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.20 |  | 0.002 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:55 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:59 PM | 07.02.2025 12:01 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506641307 | A096996 | 2 | 3 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.99 |  | 0.014 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:55 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:59 PM | 07.02.2025 12:01 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506641308 | A096996 | 3 | 3 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.27 |  | 0.006 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:55 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:59 PM | 07.02.2025 12:01 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506641309 | A096996 | 4 | 3 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.70 |  | 0.283 |  |  |  | UNDILUTED | 29.80 | 07.02.2025 01:55 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:59 PM | 07.02.2025 12:01 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506641310 | A096996 | 5 | 3 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.67 |  | 0.002 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:55 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:59 PM | 07.02.2025 12:01 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506641311 | A096996 | 6 | 3 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 131,231 Copies/mL |  | 15.57 |  | 0.390 |  |  |  | UNDILUTED | 18.81 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:59 PM | 07.02.2025 12:01 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641312 | A096996 | 7 | 3 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.94 |  | 0.315 | Target Not Detected |  |  | UNDILUTED | 28.36 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:59 PM | 07.02.2025 12:01 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641313 | A096996 | 8 | 3 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.74 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:59 PM | 07.02.2025 12:01 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641314 | A096996 | 9 | 3 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.84 |  | 0.234 | Target Not Detected |  |  | UNDILUTED | 29.46 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:59 PM | 07.02.2025 12:01 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641315 | A096996 | 10 | 4 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.67 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:59 PM | 07.02.2025 12:17 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641316 | A096996 | 11 | 4 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.24 |  | 0.008 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:59 PM | 07.02.2025 12:17 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641317 | A075890 | 4 | 1 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.13 |  | 0.002 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 11:15 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:21 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641318 | A075890 | 5 | 1 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.09 |  | 0.001 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:15 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:21 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641319 | A075890 | 6 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.96 |  | 0.273 | Target Not Detected |  |  | UNDILUTED | 30.72 | 07.01.2025 11:15 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:21 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641320 | A075890 | 7 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.40 |  | 0.257 |  |  |  | UNDILUTED | 29.66 | 07.01.2025 11:15 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:21 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641401 | A075890 | 8 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.40 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:15 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:21 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641402 | A075890 | 9 | 1 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.22 |  | 0.002 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 11:15 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:21 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641403 | A075890 | 10 | 2 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.75 |  | 0.202 | Target Not Detected |  |  | UNDILUTED | 30.19 | 07.01.2025 11:31 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641404 | A075890 | 11 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.00 |  | 0.251 | Target Not Detected |  |  | UNDILUTED | 27.50 | 07.01.2025 11:31 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641405 | A075890 | 12 | 2 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.17 |  | 0.004 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 11:31 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641406 | A096998 | 1 | 4 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.40 |  | 0.276 | Target Not Detected |  |  | UNDILUTED | 26.31 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:17 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641407 | A096998 | 2 | 4 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.66 |  | 0.234 | Target Not Detected |  |  | UNDILUTED | 30.28 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:17 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641408 | A096998 | 3 | 4 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.16 |  | 0.348 |  |  |  | UNDILUTED | 26.62 | 07.02.2025 02:11 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:17 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506641409 | A096998 | 4 | 4 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.26 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:17 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641410 | A096998 | 5 | 4 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.68 |  | 0.130 | Target Not Detected |  |  | UNDILUTED | 32.48 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:17 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641411 | A096998 | 6 | 4 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.91 |  | 0.368 | Target Not Detected |  |  | UNDILUTED | 27.82 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:17 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641412 | A096998 | 7 | 4 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.35 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 02:11 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00374,00976 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 09:00 PM | 07.02.2025 12:17 AM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506641413 | A075981 | 12 | 4 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.62 |  | 0.339 | Target Not Detected |  |  | UNDILUTED | 28.49 | 07.02.2025 12:03 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641414 | A076014 | 1 | 2 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.53 |  | 0.235 | Target Not Detected |  |  | UNDILUTED | 27.81 | 07.01.2025 11:31 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641415 | A076014 | 2 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 09:38 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 5011 - Pipettor sample aspiration error type: (Empty Tube Detected), pipettor: (1) | 399437 | 03505,00826 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 08:55 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506641416 | A076014 | 3 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.80 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:31 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641417 | A076014 | 4 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 09:38 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 5011 - Pipettor sample aspiration error type: (Empty Tube Detected), pipettor: (3) | 399437 | 03505,00826 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 08:55 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506641418 | A076014 | 5 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.71 |  | 0.001 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:31 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641419 | A076014 | 6 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.55 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:31 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641420 | A076014 | 7 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.11 |  | 0.061 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 11:31 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641501 | A076014 | 8 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.94 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:31 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641502 | A076014 | 9 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.80 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:31 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641503 | A076014 | 10 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.59 |  | 0.261 | Target Not Detected |  |  | UNDILUTED | 26.72 | 07.01.2025 11:47 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641504 | A076014 | 11 | 3 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.64 |  | 0.008 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:47 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641505 | A076014 | 12 | 3 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.50 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:47 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641506 | A075890 | 1 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.42 |  | 0.221 | Target Not Detected |  |  | UNDILUTED | 30.42 | 07.01.2025 11:15 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00456,01439 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:21 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641507 | A075890 | 2 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.89 |  | 0.187 | Target Not Detected |  |  | UNDILUTED | 29.10 | 07.01.2025 11:15 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:21 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641508 | A075890 | 3 | 1 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.92 |  | 0.209 | Target Not Detected |  |  | UNDILUTED | 28.78 | 07.01.2025 11:15 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:54 PM | 07.01.2025 09:21 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641509 | A075772 | 8 | 4 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.23 |  | 0.006 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:03 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641510 | A075772 | 9 | 4 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.53 |  | 0.363 | Target Not Detected |  |  | UNDILUTED | 29.65 | 07.02.2025 12:03 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641511 | A075772 | 10 | 1 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.58 |  | 0.264 | Target Not Detected |  |  | UNDILUTED | 28.13 | 07.02.2025 12:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641512 | A075772 | 11 | 1 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.59 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641513 | A075772 | 12 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.43 |  | 0.245 | Target Not Detected |  |  | UNDILUTED | 27.98 | 07.02.2025 12:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641514 | A075981 | 1 | 3 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.19 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 11:47 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641515 | A075981 | 2 | 3 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 |  |  | 16.15 |  | 0.384 |  |  |  | UNDILUTED | 24.34 | 07.01.2025 11:47 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641516 | A075981 | 3 | 3 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.81 |  | 0.310 | Target Not Detected |  |  | UNDILUTED | 29.61 | 07.01.2025 11:47 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641517 | A075981 | 4 | 3 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.69 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:47 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641518 | A075981 | 5 | 3 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.59 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 11:47 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641519 | A075981 | 6 | 3 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 09:54 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 5011 - Pipettor sample aspiration error type: (Empty Tube Detected), pipettor: (1) | 399437 | 03505,00826 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506641520 | A075981 | 7 | 3 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.05 |  | 0.289 | Target Not Detected |  |  | UNDILUTED | 28.75 | 07.01.2025 11:47 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641601 | A075981 | 8 | 3 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.81 |  | 0.384 | Target Not Detected |  |  | UNDILUTED | 26.86 | 07.01.2025 11:47 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641602 | A075981 | 9 | 3 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.52 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:47 PM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641603 | A075981 | 10 | 4 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.57 |  | 0.278 | Target Not Detected |  |  | UNDILUTED | 28.18 | 07.02.2025 12:03 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641604 | A075981 | 11 | 4 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.89 |  | 0.284 | Target Not Detected |  |  | UNDILUTED | 27.51 | 07.02.2025 12:03 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:55 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641605 | A096571 | 4 | 1 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.31 |  | 0.244 |  |  |  | UNDILUTED | 28.80 | 07.02.2025 12:19 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641606 | A096571 | 5 | 1 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.86 |  | 0.216 | Target Not Detected |  |  | UNDILUTED | 28.94 | 07.02.2025 12:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641607 | A096571 | 6 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.74 |  | 0.279 |  |  |  | UNDILUTED | 26.83 | 07.02.2025 12:19 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641608 | A096571 | 7 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.04 |  | 0.256 | Target Not Detected |  |  | UNDILUTED | 28.47 | 07.02.2025 12:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641609 | A096571 | 8 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.17 |  | 0.005 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:19 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641610 | A096571 | 9 | 1 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.05 |  | 0.186 | Target Not Detected |  |  | UNDILUTED | 29.62 | 07.02.2025 12:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641611 | A096571 | 10 | 2 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.76 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641612 | A096571 | 11 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.33 |  | 0.011 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:35 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641613 | A096571 | 12 | 2 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.37 |  | 0.223 |  |  |  | UNDILUTED | 29.10 | 07.02.2025 12:35 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641614 | A075772 | 1 | 4 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.14 |  | 0.253 | Target Not Detected |  |  | UNDILUTED | 28.70 | 07.02.2025 12:03 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641615 | A075772 | 2 | 4 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.91 |  | 0.186 | Target Not Detected |  |  | UNDILUTED | 31.06 | 07.02.2025 12:03 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641616 | A075772 | 3 | 4 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:03 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 5013 - Liquid level detection error code (70) on pipettor (2). | 399437 | 03505,00826 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 08:56 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506641617 | A075772 | 4 | 4 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:03 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 5013 - Liquid level detection error code (70) on pipettor (3). | 399437 | 03505,00826 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 08:56 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506641618 | A075772 | 5 | 4 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:03 PM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 5013 - Liquid level detection error code (70) on pipettor (4). | 399437 | 03505,00826 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 08:56 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506641619 | A075772 | 6 | 4 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.57 |  | 0.009 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:03 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641620 | A075772 | 7 | 4 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.83 |  | 0.188 | Target Not Detected |  |  | UNDILUTED | 31.77 | 07.02.2025 12:03 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:10 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641701 | A075884 | 12 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.16 |  | 0.385 | Target Not Detected |  |  | UNDILUTED | 28.28 | 07.02.2025 12:37 AM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 03639,00836 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 08:50 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 02624,52667 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641702 | A075976 | 1 | 2 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.76 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641703 | A075976 | 2 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.49 |  | 0.230 | Target Not Detected |  |  | UNDILUTED | 28.95 | 07.02.2025 12:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641704 | A075976 | 3 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.02 |  | 0.218 | Target Not Detected |  |  | UNDILUTED | 31.25 | 07.02.2025 12:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641705 | A075976 | 4 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.55 |  | 0.278 | Target Not Detected |  |  | UNDILUTED | 28.84 | 07.02.2025 12:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641706 | A075976 | 5 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.72 |  | 0.284 | Target Not Detected |  |  | UNDILUTED | 29.25 | 07.02.2025 12:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641707 | A075976 | 6 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.06 |  | 0.265 | Target Not Detected |  |  | UNDILUTED | 30.76 | 07.02.2025 12:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641708 | A075976 | 7 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.92 |  | 0.297 | Target Not Detected |  |  | UNDILUTED | 27.73 | 07.02.2025 12:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641709 | A075976 | 8 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.95 |  | 0.240 | Target Not Detected |  |  | UNDILUTED | 28.51 | 07.02.2025 12:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641710 | A075976 | 9 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.85 |  | 0.374 | Target Not Detected |  |  | UNDILUTED | 26.34 | 07.02.2025 12:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641711 | A075976 | 10 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.61 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:51 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:58 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641712 | A075976 | 11 | 3 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.67 |  | 0.255 |  |  |  | UNDILUTED | 27.59 | 07.02.2025 12:51 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:58 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641713 | A075976 | 12 | 3 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.39 |  | 0.323 |  |  |  | UNDILUTED | 28.06 | 07.02.2025 12:51 AM | Exception |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.09, 16.09]. | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:57 PM | 07.01.2025 10:58 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506641714 | A096571 | 1 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.46 |  | 0.219 | Target Not Detected |  |  | UNDILUTED | 29.96 | 07.02.2025 12:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 03505,00826 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641715 | A096571 | 2 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.50 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641716 | A096571 | 3 | 1 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.79 |  | 0.270 | Target Not Detected |  |  | UNDILUTED | 29.15 | 07.02.2025 12:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:04 AM | M00846 | 1.8.0 |  | 399437 | 00477,00813 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 07:09 PM | 01.31.2026 11:59 PM | 07.01.2025 08:56 PM | 07.01.2025 10:26 PM | 404866 | 06.10.2026 11:59 PM | 00600,51634 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096995 | 2 | Not Detected |  |  | 06.30.2025 01:55 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096995 | 4 | 635 Copies/mL |  |  | 06.30.2025 01:55 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096995 | 6 | 66,110 Copies/mL |  |  | 06.30.2025 01:55 PM
2506641717 | A075884 | 8 | 4 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.88 |  | 0.273 | Target Not Detected |  |  | UNDILUTED | 29.68 | 07.02.2025 12:21 AM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 03639,00836 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 08:50 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 02624,52667 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641718 | A075884 | 9 | 4 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.60 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:21 AM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 03639,00836 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 08:50 PM | 07.01.2025 10:27 PM | 404866 | 06.10.2026 11:59 PM | 02624,52667 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641719 | A075884 | 10 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.33 |  | 0.282 | Target Not Detected |  |  | UNDILUTED | 28.37 | 07.02.2025 12:36 AM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 03639,00836 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 08:50 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 02624,52667 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506641720 | A075884 | 11 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.54 |  | 0.300 | Target Not Detected |  |  | UNDILUTED | 26.00 | 07.02.2025 12:37 AM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 03639,00836 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 08:50 PM | 07.01.2025 10:42 PM | 404866 | 06.10.2026 11:59 PM | 02624,52667 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506642801 | A076043 | 10 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:13 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642802 | A076043 | 11 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:13 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642803 | A076043 | 12 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:14 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642804 | A075933 | 1 | 3 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.97 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:43 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:49 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642805 | A075933 | 2 | 3 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.62 |  | 0.090 |  |  |  | UNDILUTED | 29.57 | 07.02.2025 12:43 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:49 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506642806 | A075933 | 3 | 3 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:43 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:47 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642807 | A075933 | 4 | 3 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:43 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:47 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642808 | A075933 | 5 | 3 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:43 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:47 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642809 | A075933 | 6 | 3 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:43 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:47 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642810 | A075933 | 7 | 3 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:44 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:47 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642811 | A075933 | 8 | 3 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:44 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:47 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642812 | A075933 | 9 | 3 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:44 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:47 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642813 | A076044 | 4 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.77 |  | 0.219 | Target Not Detected |  |  | UNDILUTED | 27.05 | 07.02.2025 01:15 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:48 PM | 07.01.2025 11:21 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642814 | A076044 | 5 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.00 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:15 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:48 PM | 07.01.2025 11:21 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642815 | A076044 | 6 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.42 |  | 0.261 |  |  |  | UNDILUTED | 26.94 | 07.02.2025 01:15 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:48 PM | 07.01.2025 11:21 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506642816 | A076044 | 7 | 1 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.47 |  | 0.002 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:15 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:48 PM | 07.01.2025 11:21 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506642817 | A075933 | 12 | 4 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:59 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:47 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642818 | A096986 | 1 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.04 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 12:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642819 | A096986 | 2 | 2 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.91 |  | 0.181 | Target Not Detected |  |  | UNDILUTED | 27.11 | 07.02.2025 12:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642820 | A096986 | 3 | 2 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.16 |  | 0.232 | Target Not Detected |  |  | UNDILUTED | 26.57 | 07.02.2025 12:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642901 | A096986 | 4 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.61 |  | 0.006 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:27 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506642902 | A096986 | 5 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.30 |  | 0.216 | Target Not Detected |  |  | UNDILUTED | 26.60 | 07.02.2025 12:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642903 | A096986 | 6 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.41 |  | 0.257 |  |  |  | UNDILUTED | 25.02 | 07.02.2025 12:27 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506642904 | A096986 | 7 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.07 |  | 0.269 | Detected <LLOQ |  |  | UNDILUTED | 25.62 | 07.02.2025 12:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642905 | A096986 | 8 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.30 |  | 0.278 | Target Not Detected |  |  | UNDILUTED | 26.59 | 07.02.2025 12:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642906 | A096986 | 9 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 17.01 |  | 0.301 |  |  |  | UNDILUTED | 25.63 | 07.02.2025 12:27 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506642907 | A096986 | 10 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.67 |  | 0.263 |  |  |  | UNDILUTED | 26.26 | 07.02.2025 12:27 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506642908 | A096986 | 11 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.88 |  | 0.242 | Target Not Detected |  |  | UNDILUTED | 26.24 | 07.02.2025 12:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642909 | A096986 | 12 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.15 |  | 0.169 | Target Not Detected |  |  | UNDILUTED | 27.46 | 07.02.2025 12:43 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:47 PM | 07.01.2025 10:49 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642910 | A076043 | 7 | 4 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:00 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642911 | A076043 | 8 | 1 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:13 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642912 | A076043 | 9 | 1 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:13 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506642913 | A097009 | 2 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.76 |  | 0.277 |  |  |  | UNDILUTED | 27.13 | 07.02.2025 12:11 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:18 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506642914 | A097009 | 3 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.03 |  | 0.174 | Target Not Detected |  |  | UNDILUTED | 27.12 | 07.02.2025 12:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:18 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642915 | A097009 | 4 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.35 |  | 0.277 | Detected <LLOQ |  |  | UNDILUTED | 25.60 | 07.02.2025 12:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:18 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642916 | A097009 | 5 | 1 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 451 Copies/mL |  | 16.10 |  | 0.284 |  |  |  | UNDILUTED | 24.92 | 07.02.2025 12:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:18 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642917 | A097009 | 6 | 1 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.18 |  | 0.248 | Target Not Detected |  |  | UNDILUTED | 27.81 | 07.02.2025 12:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:18 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642918 | A097009 | 7 | 1 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.32 |  | 0.082 | Target Not Detected |  |  | UNDILUTED | 29.83 | 07.02.2025 12:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:18 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642919 | A097009 | 8 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 15.94 |  | 0.281 | Detected <LLOQ |  |  | UNDILUTED | 25.71 | 07.02.2025 12:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:18 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506642920 | A097009 | 9 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.61 |  | 0.169 |  |  |  | UNDILUTED | 28.66 | 07.02.2025 12:11 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:18 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643001 | A097009 | 10 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 17.48 |  | 0.337 |  |  |  | UNDILUTED | 25.36 | 07.02.2025 12:11 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:18 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643002 | A097009 | 11 | 1 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.76 |  | 0.312 |  |  |  | UNDILUTED | 26.52 | 07.02.2025 12:11 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:18 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643003 | A097009 | 12 | 2 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.54 |  | 0.002 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 12:27 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02690,00505 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:34 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643007 | A075933 | 10 | 3 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:44 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:47 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506643008 | A075933 | 11 | 4 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:57 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:47 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506643018 | A075930 | 1 | 4 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.23 |  | 0.283 | Detected <LLOQ |  |  | UNDILUTED | 25.40 | 07.01.2025 11:55 PM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:45 PM | 07.01.2025 10:02 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643019 | A075930 | 2 | 4 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.74 |  | 0.239 |  |  |  | UNDILUTED | 26.57 | 07.01.2025 11:55 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:45 PM | 07.01.2025 10:02 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643020 | A075930 | 3 | 4 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.54 |  | 0.088 |  |  |  | UNDILUTED | 28.57 | 07.01.2025 11:55 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:45 PM | 07.01.2025 10:02 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643101 | A075930 | 4 | 4 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 |  |  | 16.91 |  | 0.314 |  |  |  | UNDILUTED | 24.84 | 07.01.2025 11:55 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:45 PM | 07.01.2025 10:02 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643102 | A075930 | 5 | 4 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.17 |  | 0.309 | Target Not Detected |  |  | UNDILUTED | 26.73 | 07.01.2025 11:55 PM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:45 PM | 07.01.2025 10:02 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643103 | A075930 | 6 | 4 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 |  |  | 17.23 |  | 0.305 |  |  |  | UNDILUTED | 24.36 | 07.01.2025 11:55 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:45 PM | 07.01.2025 10:02 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643104 | A097009 | 1 | 1 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.71 |  | 0.108 |  |  |  | UNDILUTED | 29.84 | 07.02.2025 12:11 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 02734,00506 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:46 PM | 07.01.2025 10:18 PM | 404866 | 06.10.2026 11:59 PM | 01630,50261 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643105 | A097000 | 10 | 3 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.11 |  | 0.246 | Target Not Detected |  |  | UNDILUTED | 26.24 | 07.02.2025 01:47 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643106 | A097000 | 11 | 3 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 741 Copies/mL |  | 16.14 |  | 0.319 |  |  |  | UNDILUTED | 24.59 | 07.02.2025 01:47 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643107 | A097000 | 12 | 3 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 15.64 |  | 0.324 | Detected <LLOQ |  |  | UNDILUTED | 25.67 | 07.02.2025 01:47 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643108 | A076044 | 1 | 1 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:14 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506643109 | A076044 | 2 | 1 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:14 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506643110 | A076044 | 3 | 1 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:14 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506643111 | A075931 | 10 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.55 |  | 0.260 |  |  |  | UNDILUTED | 25.01 | 07.02.2025 02:18 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:50 PM | 07.02.2025 12:25 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643112 | A075931 | 11 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 3,933 Copies/mL |  | 16.24 |  | 0.349 |  |  |  | UNDILUTED | 23.07 | 07.02.2025 02:18 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:50 PM | 07.02.2025 12:25 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643113 | A075931 | 12 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 |  |  | 16.39 |  | 0.364 |  |  |  | UNDILUTED | 23.24 | 07.02.2025 02:18 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:50 PM | 07.02.2025 12:25 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643114 | A097008 | 1 | 3 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 15.83 |  | 0.332 | Detected <LLOQ |  |  | UNDILUTED | 25.00 | 07.02.2025 01:47 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643115 | A097008 | 2 | 3 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.09 |  | 0.326 | Detected <LLOQ |  |  | UNDILUTED | 25.22 | 07.02.2025 01:47 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643116 | A097008 | 3 | 3 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 950 Copies/mL |  | 15.14 |  | 0.301 |  |  |  | UNDILUTED | 24.40 | 07.02.2025 01:47 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643117 | A097008 | 4 | 3 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.77 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:47 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643118 | A097008 | 5 | 3 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.27 |  | 0.303 | Target Not Detected |  |  | UNDILUTED | 26.83 | 07.02.2025 01:47 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643119 | A097008 | 6 | 3 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.80 |  | 0.246 |  |  |  | UNDILUTED | 26.50 | 07.02.2025 01:47 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643120 | A097008 | 7 | 3 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.77 |  | 0.233 | Target Not Detected |  |  | UNDILUTED | 26.83 | 07.02.2025 01:47 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643201 | A076044 | 12 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.88 |  | 0.156 |  |  |  | UNDILUTED | 27.47 | 07.02.2025 01:31 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:48 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643202 | A076043 | 1 | 4 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:59 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506643203 | A076043 | 2 | 4 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:59 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506643204 | A076043 | 3 | 4 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 10:59 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506643205 | A076043 | 4 | 4 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:00 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506643206 | A076043 | 5 | 4 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:00 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506643207 | A076043 | 6 | 4 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | [8 empty] | UNDILUTED |  | 07.01.2025 11:00 PM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 5002 - Movement error on z drive (2). | 399437 | 02690,00505 | 07.14.2026 11:59 PM |  |  |  | 07.01.2025 09:48 PM |  | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 | [36 empty] | HIV-DBS
2506643208 | A097000 | 1 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.98 |  | 0.148 | Target Not Detected |  |  | UNDILUTED | 28.89 | 07.02.2025 01:31 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643209 | A097000 | 2 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.49 |  | 0.179 |  |  |  | UNDILUTED | 28.15 | 07.02.2025 01:32 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643210 | A097000 | 3 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.03 |  | 0.256 | Detected <LLOQ |  |  | UNDILUTED | 25.06 | 07.02.2025 01:32 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643211 | A097000 | 4 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.87 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:32 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643212 | A097000 | 5 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.53 |  | 0.195 |  |  |  | UNDILUTED | 29.47 | 07.02.2025 01:32 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643213 | A097000 | 6 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.84 |  | 0.172 | Target Not Detected |  |  | UNDILUTED | 29.03 | 07.02.2025 01:32 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643214 | A097000 | 7 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.59 |  | 0.298 | Target Not Detected |  |  | UNDILUTED | 25.82 | 07.02.2025 01:32 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643215 | A097000 | 8 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.32 |  | 0.117 | Target Not Detected |  |  | UNDILUTED | 28.41 | 07.02.2025 01:47 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643216 | A097000 | 9 | 3 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.13 |  | 0.217 | Target Not Detected |  |  | UNDILUTED | 26.87 | 07.02.2025 01:47 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:49 PM | 07.01.2025 11:53 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643217 | A076044 | 8 | 2 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.86 |  | 0.023 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:31 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:48 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643218 | A076044 | 9 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.42 |  | 0.264 |  |  |  | UNDILUTED | 26.62 | 07.02.2025 01:31 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:48 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506643219 | A076044 | 10 | 2 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 926 Copies/mL |  | 16.18 |  | 0.310 |  |  |  | UNDILUTED | 24.42 | 07.02.2025 01:31 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:48 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643220 | A076044 | 11 | 2 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 2,304 Copies/mL |  | 15.90 |  | 0.312 |  |  |  | UNDILUTED | 23.61 | 07.02.2025 01:31 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:48 PM | 07.01.2025 11:37 PM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506643801 | A075980 | 3 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.95 |  | 0.257 | Target Not Detected |  |  | UNDILUTED | 29.08 | 07.02.2025 01:39 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643802 | A075980 | 4 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.30 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:39 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643803 | A075980 | 5 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.58 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:39 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643804 | A075980 | 6 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.76 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:39 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643805 | A075980 | 7 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.95 |  | 0.266 |  |  |  | UNDILUTED | 28.57 | 07.02.2025 01:39 AM | Exception |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.48, 16.48]. | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427
2506643806 | A075980 | 8 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.93 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:39 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643807 | A075980 | 9 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.60 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:39 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643808 | A075980 | 10 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.36 |  | 0.326 | Target Not Detected |  |  | UNDILUTED | 27.59 | 07.02.2025 01:39 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643809 | A075980 | 11 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 |  |  | 16.49 |  | 0.368 |  |  |  | UNDILUTED | 24.41 | 07.02.2025 01:55 AM | Exception |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.48, 16.48]. | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427
2506643810 | A075980 | 12 | 3 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.89 |  | 0.008 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643811 | A076016 | 1 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.56 |  | 0.257 | Target Not Detected |  |  | UNDILUTED | 27.66 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643812 | A076016 | 2 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.84 |  | 0.288 | Target Not Detected |  |  | UNDILUTED | 26.24 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643813 | A076016 | 3 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.17 |  | 0.377 | Target Not Detected |  |  | UNDILUTED | 25.92 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643814 | A076016 | 4 | 1 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.56 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643815 | A076016 | 5 | 1 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.73 |  | 0.264 | Target Not Detected |  |  | UNDILUTED | 28.71 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643816 | A076016 | 6 | 1 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.86 |  | 0.173 | Target Not Detected |  |  | UNDILUTED | 31.69 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643817 | A075974 | 11 | 1 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | -1.00 |  | 0.015 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 02:27 AM | Exception |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 | 9212 - Internal control failed. | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427
2506643818 | A075974 | 12 | 1 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.94 |  | 0.359 | Target Not Detected |  |  | UNDILUTED | 26.76 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643819 | A075979 | 1 | 3 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.83 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643820 | A075979 | 2 | 3 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.52 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:55 AM | Exception |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.48, 16.48]. | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427
2506643901 | A075979 | 3 | 3 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.47 |  | 0.251 | Target Not Detected |  |  | UNDILUTED | 29.56 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643902 | A075979 | 4 | 3 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.22 |  | 0.210 | Target Not Detected |  |  | UNDILUTED | 29.75 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643903 | A075979 | 5 | 3 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.39 |  | 0.286 | Target Not Detected |  |  | UNDILUTED | 27.77 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643904 | A075979 | 6 | 3 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.76 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:55 AM | Exception |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.48, 16.48]. | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427
2506643905 | A075979 | 7 | 3 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.48 |  | 0.009 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:55 AM | Exception |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.48, 16.48]. | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427
2506643906 | A075979 | 8 | 3 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.82 |  | 0.234 | Target Not Detected |  |  | UNDILUTED | 28.75 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643907 | A075979 | 9 | 3 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.44 |  | 0.234 | Target Not Detected |  |  | UNDILUTED | 29.90 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643908 | A075979 | 10 | 3 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.60 |  | 0.217 | Target Not Detected |  |  | UNDILUTED | 31.11 | 07.02.2025 01:55 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:02 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643909 | A075979 | 11 | 4 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.00 |  | 0.252 | Target Not Detected |  |  | UNDILUTED | 27.54 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643910 | A075979 | 12 | 4 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.69 |  | 0.255 | Target Not Detected |  |  | UNDILUTED | 26.98 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643911 | A075980 | 1 | 2 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.97 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:39 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643912 | A075980 | 2 | 2 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.43 |  | 0.362 | Target Not Detected |  |  | UNDILUTED | 26.13 | 07.02.2025 01:39 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:26 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643913 | A075973 | 7 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.33 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02979,03182 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643914 | A075973 | 8 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.43 |  | 0.225 | Target Not Detected |  |  | UNDILUTED | 30.32 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02979,03182 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643915 | A075973 | 9 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.15 |  | 0.240 | Target Not Detected |  |  | UNDILUTED | 29.43 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02979,03182 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643916 | A075973 | 10 | 1 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.78 |  | 0.378 | Target Not Detected |  |  | UNDILUTED | 26.33 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02979,03182 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643917 | A075973 | 11 | 2 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 15.66 |  | 0.307 | Detected <LLOQ |  |  | UNDILUTED | 25.10 | 07.02.2025 02:43 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02979,03182 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:50 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643918 | A075973 | 12 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.90 |  | 0.262 | Target Not Detected |  |  | UNDILUTED | 27.47 | 07.02.2025 02:43 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02979,03182 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:50 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643919 | A075974 | 1 | 4 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.29 |  | 0.228 | Target Not Detected |  |  | UNDILUTED | 28.68 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506643920 | A075974 | 2 | 4 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.82 |  | 0.234 | Target Not Detected |  |  | UNDILUTED | 27.99 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644001 | A075974 | 3 | 4 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.70 |  | 0.312 | Target Not Detected |  |  | UNDILUTED | 28.65 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644002 | A075974 | 4 | 4 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.82 |  | 0.303 | Target Not Detected |  |  | UNDILUTED | 28.04 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644003 | A075974 | 5 | 4 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.98 |  | 0.281 | Target Not Detected |  |  | UNDILUTED | 27.81 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644004 | A075974 | 6 | 4 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 15.40 |  | 0.363 | Detected <LLOQ |  |  | UNDILUTED | 25.11 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644005 | A075974 | 7 | 4 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.34 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644006 | A075974 | 8 | 4 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.00 |  | 0.008 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644007 | A075974 | 9 | 4 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.82 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644008 | A075974 | 10 | 4 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.77 |  | 0.395 | Target Not Detected |  |  | UNDILUTED | 26.49 | 07.02.2025 02:11 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:18 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644009 | A145651 | 3 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.28 |  | 0.244 | Target Not Detected |  |  | UNDILUTED | 30.17 | 07.01.2025 11:32 PM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 01005,01260 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:22 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644010 | A145651 | 4 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.62 |  | 0.249 | Target Not Detected |  |  | UNDILUTED | 29.33 | 07.01.2025 11:32 PM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 01005,01260 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:22 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644011 | A145651 | 5 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 8,876 Copies/mL |  | 16.17 |  | 0.449 |  |  |  | UNDILUTED | 21.93 | 07.01.2025 11:32 PM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 01005,01260 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:22 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644012 | A145651 | 6 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.93 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:32 PM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 01005,01260 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:22 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644013 | A145651 | 7 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.70 |  | 0.314 | Target Not Detected |  |  | UNDILUTED | 29.86 | 07.01.2025 11:32 PM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 01005,01260 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:22 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644014 | A145651 | 8 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.89 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 11:32 PM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 01005,01260 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:22 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644015 | A145651 | 9 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.56 |  | 0.345 |  |  |  | UNDILUTED | 26.93 | 07.01.2025 11:32 PM | Exception |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.48, 16.48]. | 399437 | 01005,01260 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:22 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506644016 | A145651 | 10 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 15.62 |  | 0.362 | Detected <LLOQ |  |  | UNDILUTED | 25.28 | 07.01.2025 11:32 PM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 01005,01260 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:22 PM | 07.01.2025 09:38 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644017 | A145651 | 11 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.58 |  | 0.221 | Target Not Detected |  |  | UNDILUTED | 30.05 | 07.01.2025 11:48 PM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 01005,01260 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:22 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644018 | A145651 | 12 | 3 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.92 |  | 0.266 | Target Not Detected |  |  | UNDILUTED | 28.16 | 07.01.2025 11:48 PM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 01005,01260 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:22 PM | 07.01.2025 09:54 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644019 | A075973 | 1 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.81 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644020 | A075973 | 2 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.66 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644101 | A075973 | 3 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.36 |  | 0.278 | Target Not Detected |  |  | UNDILUTED | 28.37 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644102 | A075973 | 4 | 1 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.19 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644103 | A075973 | 5 | 1 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.26 |  | 0.290 | Target Not Detected |  |  | UNDILUTED | 29.52 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644104 | A075973 | 6 | 1 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.64 |  | 0.295 | Target Not Detected |  |  | UNDILUTED | 27.87 | 07.02.2025 02:27 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:27 PM | 07.02.2025 12:34 AM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644105 | A075932 | 3 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.66 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:41 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644106 | A075932 | 4 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.74 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:41 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644107 | A075932 | 5 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.39 |  | 0.011 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 02:35 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:41 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506644108 | A075932 | 6 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.12 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:35 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:41 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644109 | A075932 | 7 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.65 |  | 0.308 |  |  |  | UNDILUTED | 26.42 | 07.02.2025 02:35 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:41 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506644110 | A075932 | 8 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.81 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:50 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:55 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644111 | A075932 | 9 | 3 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.45 |  | 0.170 | Target Not Detected |  |  | UNDILUTED | 29.85 | 07.02.2025 02:50 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:55 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644112 | A075932 | 10 | 3 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.06 |  | 0.271 | Target Not Detected |  |  | UNDILUTED | 28.74 | 07.02.2025 02:50 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:55 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644113 | A075932 | 11 | 3 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.87 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:50 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:55 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644114 | A075932 | 12 | 3 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.89 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:50 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:55 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644115 | A076047 | 1 | 1 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.17 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:50 PM | 07.02.2025 12:25 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644116 | A076047 | 2 | 1 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.25 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:50 PM | 07.02.2025 12:25 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644117 | A076047 | 3 | 1 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 950 Copies/mL |  | 15.53 |  | 0.381 |  |  |  | UNDILUTED | 24.40 | 07.02.2025 02:19 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:50 PM | 07.02.2025 12:25 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644118 | A076047 | 4 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.97 |  | 0.001 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 02:19 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:50 PM | 07.02.2025 12:25 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506644119 | A076047 | 5 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.93 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 02:19 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:50 PM | 07.02.2025 12:25 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506644120 | A076047 | 6 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.54 |  | 0.008 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 02:19 AM | Exception |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.38, 16.38]. | 399437 | 01213,01015 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:50 PM | 07.02.2025 12:25 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614
2506644201 | A076016 | 11 | 2 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.68 |  | 0.242 | Target Not Detected |  |  | UNDILUTED | 28.77 | 07.02.2025 01:39 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644202 | A076016 | 12 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.99 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:39 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:46 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644203 | A075889 | 1 | 4 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.79 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:07 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:14 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644204 | A075889 | 2 | 4 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.65 |  | 0.444 | Target Not Detected |  |  | UNDILUTED | 27.28 | 07.02.2025 01:07 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:14 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644205 | A075889 | 3 | 4 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.57 |  | 0.307 | Target Not Detected |  |  | UNDILUTED | 25.95 | 07.02.2025 01:07 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:14 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644206 | A075889 | 4 | 4 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.77 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:07 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:14 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644207 | A075889 | 5 | 4 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.24 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:07 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:14 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644208 | A075889 | 6 | 4 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.39 |  | 0.235 | Target Not Detected |  |  | UNDILUTED | 29.59 | 07.02.2025 01:07 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:14 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644209 | A075889 | 7 | 4 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.88 |  | 0.326 | Target Not Detected |  |  | UNDILUTED | 27.70 | 07.02.2025 01:08 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:14 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644210 | A075889 | 8 | 4 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.99 |  | 0.396 | Target Not Detected |  |  | UNDILUTED | 28.03 | 07.02.2025 01:08 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:14 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644211 | A075889 | 9 | 4 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.92 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:08 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:14 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644212 | A075889 | 10 | 4 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.83 |  | 0.007 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:08 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:14 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644213 | A075889 | 11 | 1 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.91 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644214 | A075889 | 12 | 1 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.94 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 00948,01016 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644215 | A075932 | 1 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.86 |  | 0.006 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:34 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:41 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644216 | A075932 | 2 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.00 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 02:34 AM | Completed |  |  |  | Archived | 07.02.2025 09:01 AM | M01028 | 1.8.0 |  | 399437 | 00859,02611 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 09:51 PM | 07.02.2025 12:41 AM | 404866 | 06.10.2026 11:59 PM | 00635,51672 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1754614 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075775 | 2 | Not Detected |  |  | 06.30.2025 02:01 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075775 | 4 | 929 Copies/mL |  |  | 06.30.2025 02:01 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075775 | 6 | 66,946 Copies/mL |  |  | 06.30.2025 02:01 PM
2506644217 | A076016 | 7 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.98 |  | 0.351 | Target Not Detected |  |  | UNDILUTED | 28.47 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644218 | A076016 | 8 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.50 |  | 0.369 | Target Not Detected |  |  | UNDILUTED | 26.23 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644219 | A076016 | 9 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.69 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644220 | A076016 | 10 | 1 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.04 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:23 AM | Completed |  |  |  | Archived | 07.02.2025 09:08 AM | M00906 | 1.8.0 |  | 399437 | 02679,03168 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:05 PM | 01.31.2026 11:59 PM | 07.01.2025 09:25 PM | 07.01.2025 11:30 PM | 404866 | 06.10.2026 11:59 PM | 01748,50710 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2861427 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A096987 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A096987 | 4 | 872 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A096987 | 6 | 69,032 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644301 | A075770 | 12 | 1 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.23 |  | 0.223 | Target Not Detected |  |  | UNDILUTED | 25.82 | 07.01.2025 09:09 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01600,03704 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:16 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644302 | A145653 | 1 | 4 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.14 |  | 0.152 | Target Not Detected |  |  | UNDILUTED | 28.79 | 07.01.2025 09:08 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644303 | A145653 | 2 | 4 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.31 |  | 0.227 | Target Not Detected |  |  | UNDILUTED | 27.98 | 07.01.2025 09:08 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644304 | A145653 | 3 | 4 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.50 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:08 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644305 | A145653 | 4 | 4 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.61 |  | 0.261 | Target Not Detected |  |  | UNDILUTED | 25.89 | 07.01.2025 09:08 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644306 | A145653 | 5 | 4 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.95 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:08 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644307 | A145653 | 6 | 4 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.28 |  | 0.161 | Target Not Detected |  |  | UNDILUTED | 28.96 | 07.01.2025 09:08 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644308 | A145653 | 7 | 4 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.08 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:08 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644309 | A145653 | 8 | 4 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.62 |  | 0.012 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:08 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644310 | A145653 | 9 | 4 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.89 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:08 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644311 | A145653 | 10 | 4 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.17 |  | 0.373 | Detected <LLOQ |  |  | UNDILUTED | 24.81 | 07.01.2025 09:08 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644312 | A145653 | 11 | 4 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.48 |  | 0.234 |  |  |  | UNDILUTED | 27.92 | 07.01.2025 09:08 PM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506644313 | A145653 | 12 | 1 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.55 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:24 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:28 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644314 | A076017 | 1 | 3 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.52 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644315 | A076017 | 2 | 3 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.89 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644316 | A076017 | 3 | 3 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.82 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644317 | A076017 | 4 | 3 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.00 |  | 0.150 | Target Not Detected |  |  | UNDILUTED | 29.80 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644318 | A076017 | 5 | 3 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.04 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644319 | A076017 | 6 | 3 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.33 |  | 0.222 | Target Not Detected |  |  | UNDILUTED | 26.64 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644320 | A076017 | 7 | 3 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.88 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644401 | A076017 | 8 | 3 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.61 |  | 0.294 | Target Not Detected |  |  | UNDILUTED | 28.08 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644402 | A076017 | 9 | 3 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.04 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644403 | A076017 | 10 | 3 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.74 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644404 | A076017 | 11 | 3 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.33 |  | 0.189 | Target Not Detected |  |  | UNDILUTED | 29.01 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644405 | A076017 | 12 | 4 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.32 |  | 0.112 | Target Not Detected |  |  | UNDILUTED | 29.02 | 07.01.2025 09:08 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:37 PM | 07.01.2025 07:15 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644406 | A075883 | 1 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.32 |  | 0.008 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 08:36 PM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209
2506644407 | A075883 | 2 | 2 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.92 |  | 0.181 | Target Not Detected |  |  | UNDILUTED | 28.04 | 07.01.2025 08:36 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644408 | A075883 | 3 | 2 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.72 |  | 0.058 | Target Not Detected |  |  | UNDILUTED | 30.24 | 07.01.2025 08:37 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644409 | A075883 | 4 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.13 |  | 0.261 | Target Not Detected |  |  | UNDILUTED | 26.88 | 07.01.2025 08:37 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644410 | A075883 | 5 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.93 |  | 0.240 |  |  |  | UNDILUTED | 28.79 | 07.01.2025 08:37 PM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209
2506644411 | A075883 | 6 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.07 |  | 0.198 | Target Not Detected |  |  | UNDILUTED | 27.95 | 07.01.2025 08:37 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644412 | A075883 | 7 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.54 |  | 0.247 |  |  |  | UNDILUTED | 27.37 | 07.01.2025 08:37 PM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209
2506644413 | A075883 | 8 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.47 |  | 0.219 |  |  |  | UNDILUTED | 28.31 | 07.01.2025 08:37 PM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209
2506644414 | A075883 | 9 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.35 |  | 0.278 |  |  |  | UNDILUTED | 25.94 | 07.01.2025 08:37 PM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209
2506644415 | A075883 | 10 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 3,206 Copies/mL |  | 15.82 |  | 0.393 |  |  |  | UNDILUTED | 22.90 | 07.01.2025 08:37 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644416 | A075883 | 11 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.58 |  | 0.339 |  |  |  | UNDILUTED | 25.75 | 07.01.2025 08:37 PM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209
2506644417 | A075883 | 12 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 15.72 |  | 0.351 | Detected <LLOQ |  |  | UNDILUTED | 24.71 | 07.01.2025 08:52 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:59 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644418 | A076015 | 1 | 1 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.51 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 08:20 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03585,01493 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:27 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644419 | A076015 | 2 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.93 |  | 0.113 | Target Not Detected |  |  | UNDILUTED | 28.85 | 07.01.2025 08:20 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03585,01493 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:27 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644420 | A076015 | 3 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.74 |  | 0.114 | Target Not Detected |  |  | UNDILUTED | 30.03 | 07.01.2025 08:20 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03585,01493 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:27 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644501 | A076015 | 4 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.99 |  | 0.178 | Target Not Detected |  |  | UNDILUTED | 28.03 | 07.01.2025 08:20 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03585,01493 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:27 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644502 | A076015 | 5 | 1 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.50 |  | 0.199 | Target Not Detected |  |  | UNDILUTED | 27.55 | 07.01.2025 08:20 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03585,01493 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:27 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644503 | A076015 | 6 | 1 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.55 |  | 0.228 | Target Not Detected |  |  | UNDILUTED | 26.67 | 07.01.2025 08:20 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03585,01493 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:27 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644504 | A076015 | 7 | 1 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.07 |  | 0.219 | Target Not Detected |  |  | UNDILUTED | 26.73 | 07.01.2025 08:21 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03585,01493 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:27 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644505 | A076015 | 8 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.81 |  | 0.236 | Target Not Detected |  |  | UNDILUTED | 28.08 | 07.01.2025 08:21 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03585,01493 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:27 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644506 | A076015 | 9 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.41 |  | 0.269 |  |  |  | UNDILUTED | 27.79 | 07.01.2025 08:21 PM | Exception |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.31, 16.31]. | 399437 | 03585,01493 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:27 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209
2506644507 | A076015 | 10 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.78 |  | 0.276 | Target Not Detected |  |  | UNDILUTED | 26.78 | 07.01.2025 08:21 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03585,01493 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:27 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644508 | A076015 | 11 | 1 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 15.57 |  | 0.288 | Detected <LLOQ |  |  | UNDILUTED | 25.33 | 07.01.2025 08:21 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03585,01493 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:27 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644509 | A076015 | 12 | 2 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.90 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 08:36 PM | Completed |  |  |  | Archived | 07.02.2025 09:09 AM | M00858 | 1.8.0 |  | 399437 | 03314,02612 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 04:04 PM | 01.31.2026 11:59 PM | 07.01.2025 05:36 PM | 07.01.2025 06:43 PM | 404866 | 06.10.2026 11:59 PM | 01755,50686 | 990020149 | 08.06.2026 11:59 PM |  |  | 2930792 | 03.25.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 801209 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A075766 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A075766 | 4 | 754 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A075766 | 6 | 56,109 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644510 | A145645 | 1 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.64 |  | 0.248 | Target Not Detected |  |  | UNDILUTED | 27.29 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644511 | A145645 | 2 | 2 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.80 |  | 0.162 | Target Not Detected |  |  | UNDILUTED | 28.13 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644512 | A145645 | 3 | 2 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.75 |  | 0.253 | Target Not Detected |  |  | UNDILUTED | 26.93 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644513 | A145645 | 4 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.85 |  | 0.198 | Target Not Detected |  |  | UNDILUTED | 29.03 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644514 | A145645 | 5 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.57 |  | 0.193 | Target Not Detected |  |  | UNDILUTED | 28.55 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644515 | A145645 | 6 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.03 |  | 0.007 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644516 | A145645 | 7 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.90 |  | 0.152 | Target Not Detected |  |  | UNDILUTED | 29.49 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644517 | A145645 | 8 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.37 |  | 0.267 | Target Not Detected |  |  | UNDILUTED | 28.22 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644518 | A145645 | 9 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.06 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644519 | A145645 | 10 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.92 |  | 0.170 | Target Not Detected |  |  | UNDILUTED | 28.78 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644520 | A145645 | 11 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.21 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644601 | A145645 | 12 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.67 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:41 PM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01317,01545 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 05:38 PM | 07.01.2025 07:45 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644602 | A075880 | 10 | 4 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 1,050 Copies/mL |  | 16.18 |  | 0.325 |  |  |  | UNDILUTED | 24.13 | 07.01.2025 10:13 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:19 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644603 | A075880 | 11 | 4 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.05 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 10:13 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:19 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644604 | A075880 | 12 | 4 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.61 |  | 0.144 | Target Not Detected |  |  | UNDILUTED | 29.22 | 07.01.2025 10:13 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:19 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644605 | A075880 | 7 | 4 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.33 |  | 0.321 | Target Not Detected |  |  | UNDILUTED | 25.63 | 07.01.2025 10:13 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:19 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644606 | A075880 | 8 | 4 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 15.06 |  | 0.333 | Detected <LLOQ |  |  | UNDILUTED | 25.29 | 07.01.2025 10:13 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:19 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644607 | A075880 | 9 | 4 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.91 |  | 0.277 | Target Not Detected |  |  | UNDILUTED | 28.13 | 07.01.2025 10:13 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:19 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644608 | A075885 | 1 | 2 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.99 |  | 0.326 | Target Not Detected |  |  | UNDILUTED | 26.96 | 07.01.2025 09:41 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644609 | A075885 | 2 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.92 |  | 0.348 |  |  |  | UNDILUTED | 25.07 | 07.01.2025 09:41 PM | Exception |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.46, 16.46]. | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506644610 | A075885 | 3 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.22 |  | 0.306 | Target Not Detected |  |  | UNDILUTED | 26.19 | 07.01.2025 09:41 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644611 | A075885 | 4 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 498 Copies/mL |  | 15.48 |  | 0.297 |  |  |  | UNDILUTED | 24.66 | 07.01.2025 09:41 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644612 | A075885 | 5 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.81 |  | 0.283 | Target Not Detected |  |  | UNDILUTED | 26.44 | 07.01.2025 09:57 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644613 | A075885 | 6 | 3 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.03 |  | 0.338 | Detected <LLOQ |  |  | UNDILUTED | 25.28 | 07.01.2025 09:57 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644614 | A075885 | 7 | 3 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.52 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 09:57 PM | Exception |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.46, 16.46]. | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506644615 | A075885 | 8 | 3 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.45 |  | 0.216 | Target Not Detected |  |  | UNDILUTED | 27.35 | 07.01.2025 09:57 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644616 | A075885 | 9 | 3 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.37 |  | 0.320 | Target Not Detected |  |  | UNDILUTED | 25.76 | 07.01.2025 09:57 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644617 | A075885 | 10 | 3 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.25 |  | 0.170 | Target Not Detected |  |  | UNDILUTED | 29.58 | 07.01.2025 09:57 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644618 | A075885 | 11 | 3 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.64 |  | 0.115 |  |  |  | UNDILUTED | 29.39 | 07.01.2025 09:57 PM | Exception |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.46, 16.46]. | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506644619 | A075885 | 12 | 3 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.64 |  | 0.201 | Target Not Detected |  |  | UNDILUTED | 28.54 | 07.01.2025 09:57 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644620 | A145279 | 1 | 4 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.10 |  | 0.004 | Target Not Detected | ME |  | UNDILUTED | -1.00 | 07.01.2025 10:13 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:44 PM | 07.01.2025 08:19 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644701 | A145279 | 2 | 4 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.62 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 10:13 PM | Exception |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.46, 16.46]. | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:44 PM | 07.01.2025 08:19 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506644702 | A145279 | 3 | 4 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.27 |  | 0.003 | Target Not Detected | ME |  | UNDILUTED | -1.00 | 07.01.2025 10:13 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:44 PM | 07.01.2025 08:19 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644703 | A075887 | 1 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.56 |  | 0.358 | Target Not Detected |  |  | UNDILUTED | 25.91 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644704 | A075887 | 2 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.78 |  | 0.242 | Target Not Detected |  |  | UNDILUTED | 27.10 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644705 | A075887 | 3 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.68 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644706 | A075887 | 4 | 1 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 1,621 Copies/mL |  | 15.63 |  | 0.355 |  |  |  | UNDILUTED | 23.76 | 07.01.2025 09:25 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:32 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644707 | A075887 | 5 | 2 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.84 |  | 0.229 | Target Not Detected |  |  | UNDILUTED | 27.03 | 07.01.2025 09:41 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644708 | A075887 | 6 | 2 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.47 |  | 0.004 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 09:41 PM | Exception |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.46, 16.46]. | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506644709 | A075887 | 7 | 2 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.84 |  | 0.102 | Target Not Detected |  |  | UNDILUTED | 29.04 | 07.01.2025 09:41 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644710 | A075887 | 8 | 2 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 15.86 |  | 0.268 | Detected <LLOQ |  |  | UNDILUTED | 25.37 | 07.01.2025 09:41 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644711 | A075887 | 9 | 2 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 17.76 |  | 0.285 |  |  |  | UNDILUTED | 26.58 | 07.01.2025 09:41 PM | Exception |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.46, 16.46]. | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506644712 | A075887 | 10 | 2 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | < 400 Copies/mL |  | 16.03 |  | 0.334 | Detected <LLOQ |  |  | UNDILUTED | 25.09 | 07.01.2025 09:41 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644713 | A075887 | 11 | 2 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.60 |  | 0.004 |  |  |  | UNDILUTED | -1.00 | 07.01.2025 09:41 PM | Exception |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.46, 16.46]. | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506644714 | A075887 | 12 | 2 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.96 |  | 0.222 | Target Not Detected |  |  | UNDILUTED | 27.42 | 07.01.2025 09:41 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:42 PM | 07.01.2025 07:47 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644715 | A075880 | 4 | 3 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.30 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 09:57 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644716 | A075880 | 5 | 4 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | 686 Copies/mL |  | 15.62 |  | 0.296 |  |  |  | UNDILUTED | 24.45 | 07.01.2025 10:13 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:19 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644717 | A075880 | 6 | 4 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.46 |  | 0.005 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.01.2025 10:13 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:19 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644718 | A075880 | 1 | 3 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.57 |  | 0.134 | Target Not Detected |  |  | UNDILUTED | 29.18 | 07.01.2025 09:57 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644719 | A075880 | 2 | 3 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.38 |  | 0.273 | Target Not Detected |  |  | UNDILUTED | 27.54 | 07.01.2025 09:57 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644720 | A075880 | 3 | 3 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.94 |  | 0.315 | Target Not Detected |  |  | UNDILUTED | 25.79 | 07.01.2025 09:57 PM | Completed |  |  |  | Archived | 07.02.2025 09:07 AM | M00833 | 1.8.0 |  | 399437 | 02108,01540 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 06:51 PM | 01.31.2026 11:59 PM | 07.01.2025 05:43 PM | 07.01.2025 08:03 PM | 404866 | 06.10.2026 11:59 PM | 02625,52681 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076048 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076048 | 4 | 824 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076048 | 6 | 73,902 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644801 | A145648 | 9 | 2 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.75 |  | 0.006 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:41 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:47 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963
2506644802 | A145648 | 10 | 2 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.10 |  | 0.279 | Target Not Detected |  |  | UNDILUTED | 28.04 | 07.02.2025 01:41 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:47 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644803 | A145648 | 11 | 2 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.95 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:41 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:47 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644804 | A145648 | 12 | 3 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.44 |  | 0.165 | Target Not Detected |  |  | UNDILUTED | 31.15 | 07.02.2025 01:56 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.02.2025 12:03 AM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644805 | A145647 | 1 | 1 | 2 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.70 |  | 0.287 | Target Not Detected |  |  | UNDILUTED | 29.04 | 07.02.2025 01:25 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644806 | A145647 | 2 | 1 | 3 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.81 |  | 0.004 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:25 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644807 | A145647 | 3 | 1 | 4 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.62 |  | 0.271 | Target Not Detected |  |  | UNDILUTED | 27.74 | 07.02.2025 01:25 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644808 | A145647 | 4 | 1 | 5 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.07 |  | 0.204 | Target Not Detected |  |  | UNDILUTED | 30.45 | 07.02.2025 01:25 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644809 | A145647 | 5 | 1 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.98 |  | 0.310 | Target Not Detected |  |  | UNDILUTED | 28.01 | 07.02.2025 01:25 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644810 | A145647 | 6 | 1 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.63 |  | 0.003 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:25 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506644811 | A145647 | 7 | 1 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.64 |  | 0.301 | Target Not Detected |  |  | UNDILUTED | 28.02 | 07.02.2025 01:25 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644812 | A145647 | 8 | 1 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.82 |  | 0.006 |  |  |  | UNDILUTED | -1.00 | 07.02.2025 01:25 AM | Exception |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 | 9210 - Internal control cycle number is too high. Valid range is [12.59, 16.59]. | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175
2506644813 | A145647 | 9 | 1 | 10 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.29 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:25 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644814 | A145647 | 10 | 1 | 11 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.91 |  | 0.329 | Target Not Detected |  |  | UNDILUTED | 28.68 | 07.02.2025 01:25 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644815 | A145647 | 11 | 1 | 12 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.31 |  | 0.305 | Target Not Detected |  |  | UNDILUTED | 28.28 | 07.02.2025 01:25 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:31 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644816 | A145647 | 12 | 2 | 1 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.00 |  | 0.220 | Target Not Detected |  |  | UNDILUTED | 29.27 | 07.02.2025 01:40 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01033,02637 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:08 PM | 07.01.2025 11:47 PM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 1720175 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644817 | A096574 | 5 | 3 | 6 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 15.64 |  | 0.307 | Target Not Detected |  |  | UNDILUTED | 29.43 | 07.02.2025 01:57 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01050,02644 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:09 PM | 07.02.2025 12:03 AM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644818 | A096574 | 6 | 3 | 7 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.19 |  | 0.003 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:57 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01050,02644 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:09 PM | 07.02.2025 12:03 AM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644819 | A096574 | 7 | 3 | 8 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.03 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:57 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01050,02644 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:09 PM | 07.02.2025 12:03 AM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM
2506644820 | A096574 | 8 | 3 | 9 | Specimen | Dried Blood Spot | Admin |  | HIV-DBS | 6200 | 1.00 | Not Detected |  | 16.18 |  | 0.002 | Target Not Detected |  |  | UNDILUTED | -1.00 | 07.02.2025 01:57 AM | Completed |  |  |  | Archived | 07.02.2025 09:11 AM | M00847 | 1.8.0 |  | 399437 | 01050,02644 | 07.14.2026 11:59 PM | 399132 | 04.03.2025 03:49 PM | 01.31.2026 11:59 PM | 07.01.2025 09:09 PM | 07.02.2025 12:03 AM | 404866 | 06.10.2026 11:59 PM | 00839,50997 | 990020149 | 08.06.2026 11:59 PM |  |  | 2471612 | 03.12.2026 11:59 PM | 990020195 | 01.22.2027 11:59 PM |  |  |  | 2344963 |  |  |  | 401574 | 04.10.2026 11:59 PM | 401572 | 04.10.2026 11:59 PM | HIV-1 NEG CTRL | A076049 | 2 | Not Detected |  |  | 06.30.2025 01:56 PM | 401584 | 04.15.2026 11:59 PM | HIV-1 LOW POS CTRL | A076049 | 4 | 777 Copies/mL |  |  | 06.30.2025 01:56 PM | 401579 | 04.15.2026 11:59 PM | HIV-1 HIGH POS CTRL | A076049 | 6 | 63,413 Copies/mL |  |  | 06.30.2025 01:56 PM